Top of Form

Generic Risk Assessment
	Risk Assessment Record



	Risk Assessment Title: Safe Use of Cryogenic Materials


	Date Produced: March 2022

	Review Date: March 2025


	Overview/Description of Activity: Activities include decanting from bulk storage containers to smaller portable containers, moving containers into and within labs, smaller scale decanting within labs and general storage e.g. of cell cultures. Typical cryogenic liquids used are nitrogen and argon. 

	Duration/Frequency of Activity: Activities such as filling/decanting tend to be on a daily basis of a short duration. Dewar’s, such as those for liquid nitrogen, are present in labs on a permanent basis.

	Location of Activity: All laboratories on University property that use cryogenic materials.
	Generic or Specific Assessment: Generic assessment to be used as basis for generating area specific risk assessments. This assessment only covers the hazards associated with the cryogenic nature of the materials not other incidental hazards such as flammability, toxicity or manual handling of containers.



	#
	Hazard(s) identified
	Who might be affected
and how
	Existing controls & measures
	Severity (a)
	Likelihood (b)
	Risk Rating 
(a x b)
	Additional control/action required

	1
	Oxygen/air displacement in lab due to a release/spill of inert cryogenic vapour while decanting/filling potentially resulting in diminished performance, fainting and asphyxiation.   


	Health effects and potential asphyxiation to person/s present within laboratory
	· External storage and decanting should be the first choice for cryogenic materials particularly for bulk storage containers

· When external storage/use is not possible then these materials must be used in a well-ventilated area (ventilation and space should be sufficient to ensure oxygen levels do not go below 19.5% under normal conditions)

· Decanting of cryogenic liquids should not be carried out in side rooms off labs (HSE action)

· Cryogenic materials should not be stored in small, enclosed spaces such cold rooms

· Only trained persons to carry out bulk decanting/filling operations

· Vessels should not be overfilled

· Number of people present should be minimised during bilk decanting/filling
· Vessels should not be left unattended while filling
	3
	2
	6
	

	2
	Release of cryogenic vapour into laboratory    

due to failure of equipment 

	Diminished performance, fainting and potential asphyxiation to person/s present within laboratory
	· Only equipment purpose designed for cryogenic materials should be used. This includes transfer tubes. Thermos flasks or similar are not suitable and may fail

· Equipment must be used in accordance with manufacturer’s instructions

· Check equipment for any defects and ice plugs, if defect is observed then equipment must be taken out of service and replaced

· Immediately re-cap any container to prevent atmospheric moisture from entering and forming an ice plug in the opening

· Keep vessels away from heat sources
	5
	2
	10
	· Ensure the requirement for the correct equipment to be used is communicated to all relevant staff

· Include in department inspections a check that the standard is being met

	3
	Rapid expansion of liquid gas leads to over pressurisation of vessel and potential explosion


	Physical impact injury to person/s present within laboratory 
	· Only equipment purpose designed for cryogenic materials should be used with safety devices such as pressure relief valves

· Equipment must be used in accordance with manufacturer’s instructions

· Vessels should never be left open but recapped immediately to prevent moisture ingress

· Vessels must be checked regularly for ice plugs which should be removed with care if observed

· Vessels should be kept away from heat sources
	5
	1
	5
	

	4
	Contact with cryogenic liquid, e.g. splash when decanting/filling 


	Person carrying out filling/decanting operation and others in immediate vicinity suffer frostbite/cold burn
	· To minimise splashing, fill slowly while the vessel cools (increased boil-off to start). 

· Do not immerse hot materials into cryogenic liquid.
· Only use containers suitable for the material being used, i.e., purpose-designed Dewar. 

· Use a transfer tube or funnel to decant liquid into narrow-necked open Dewars. 

· Use a funnel to pour liquid into narrow-necked openings in cryostats and cold traps. 

· Do not immerse hands into cryogenic liquid even when wearing gloves.
· Use tongs to place or retrieve items from cryogenic liquid. 
· When removing cell-line cages from storage, use a hook to find the handle and raise the cage.
· Wear appropriate PPE; dry thermally insulated gloves and safety goggles as a minimum. For bulk volume decanting a face shield should be worn.
	3
	2
	6
	· Ensure emergency procedures are in place to deal with an incident promptly

	5
	Contact with solid cryogenic material, e.g. cardice, when filing/emptying containers
	Person carrying out filling/emptying operation suffer frostbite/cold burn
	· Never handle cryogenic materials with bare hands

· Wear dry, thermally insulated gloves

· Use a tool such as a scoop for moving the material
	3
	1
	3
	· 

	6
	Release of cryogenic vapour when transporting in lift, e.g., due to equipment failure


	Potential asphyxiation hazard and/or cold burns if person/s present in lift
	· Cryogen liquid containers must not be accompanied in lifts. 

· If a lift is used, then it will be closed to all passengers with signage posted and preferably use of barrier tape. 

· The vessel will be manoeuvred into the lift, and the lift called to the destination floor.

	5
	2
	10
	· Ensure that the prohibition of accompanying cryogenic liquid containers in lifts is communicated to all relevant staff either through induction training and/or department handbooks 


	Assessor signature:


	Print name:


	Date:




Oxygen Depletion Calculator
[image: image1.emf]Length (metres) 0.0

Width (metres) 0.0

Height (metres) 0.0

Room Volume (l*w*h) 0.0

Volume of Liquid Nitrogen (litres) 0.00 The volume of liquid NITROGEN involved

Volume of gas produced (Vol

l

*682.5) (litres) 0.00

[liquid volume] times [expansion factor] (ref  data sheet)

Volume of gas produced (metre

3

)(Vol

gas

/1000) 0.000 volume converted to cubic metres

Volume of Liquid Argon (litres) 0.00 The volume of liquid ARGON involved

Volume of gas produced (Vol*824.3) (litres) 0.00 [liquid volume] times [expansion factor] (ref data sheet)

Volume of gas produced (metre

3

)(Vol

gas

/1000) 0.000 volume converted to cubic metres

Vol

0 

0.2095([Room Vol]-[Vol gas produced] 0.00

Volume of oxygen = [fraction of oxygen in air] times ([the 

volume of the room] - [ volume of nitrogen liberated])

%O

2

 (100*[Vol

O

]/[Room vol] #DIV/0! Volume of oxygen divided by the room volume x 100

Vol

0 

0.2095([Room Vol]-[Vol gas produced] 0.00

Volume of oxygen = [fraction of oxygen in air] times ([the 

volume of the room] - [ volume of argon liberated])

%O

2

 (100*[Vol

O

]/[Room vol] #DIV/0! Volume of oxygen divided by the room volume x 100

The calculation of percentage of oxygen in the air after the evaporation of a volume of liquefied 

nitrogen or argon gas

The room dimensions



Volume of gas produced by evaporating liquid Argon

Note: 

This assumes that there is no dilution due to fresh air being pumped into the room and does not allow for the volume 

of any furniture or people in the room

Instructions:

 Fill in the white boxes with the appropriate figures and read the percentage of oxygen remaining in the room in 

the red box.

Percentage oxygen in room for ARGON evaporation

Percentage oxygen in room for NITROGEN evaporation

Room volume

Volume of gas produced by evaporating liquid Nitrogen


Reference: based on calculations provided by BCGA and BOC
Note: Right click on spreadsheet, select worksheet object and then edit to use the calculator.

Oxygen Deficiency Symptoms
	Oxygen Content (vol %)
	Effects and Symptoms

	18-21
	No discernible symptoms can be detected

	11-18
	Diminution of physical and intellectual performance without person's knowledge

	8-11
	Possibility of fainting after a short period without prior warning

	6-8
	Fainting within a few minutes, resuscitation possible if carried out immediately

	0-6
	Fainting almost immediate, death ensues, brain damage if rescued


Note: Information taken from reputable gas association sources
� EMBED Excel.Sheet.12  ���








8                                                                                                                                                                                                            22-03-18 UHSC Safe Use of Cryogenic Materials Standard.doc

[image: image2.emf]Length (metres) 0.0

Width (metres) 0.0

Height (metres) 0.0

Room Volume (l*w*h) 0.0

Volume of Liquid Nitrogen (litres) 0.00 The volume of liquid NITROGEN involved

Volume of gas produced (Vol

l

*682.5) (litres) 0.00

[liquid volume] times [expansion factor] (ref  data sheet)

Volume of gas produced (metre

3

)(Vol

gas

/1000) 0.000 volume converted to cubic metres

Volume of Liquid Argon (litres) 0.00 The volume of liquid ARGON involved

Volume of gas produced (Vol*824.3) (litres) 0.00 [liquid volume] times [expansion factor] (ref data sheet)

Volume of gas produced (metre

3

)(Vol

gas

/1000) 0.000 volume converted to cubic metres

Vol

0 

0.2095([Room Vol]-[Vol gas produced] 0.00

Volume of oxygen = [fraction of oxygen in air] times ([the 

volume of the room] - [ volume of nitrogen liberated])

%O

2

 (100*[Vol

O

]/[Room vol] #DIV/0! Volume of oxygen divided by the room volume x 100

Vol

0 

0.2095([Room Vol]-[Vol gas produced] 0.00

Volume of oxygen = [fraction of oxygen in air] times ([the 

volume of the room] - [ volume of argon liberated])

%O

2

 (100*[Vol

O

]/[Room vol] #DIV/0! Volume of oxygen divided by the room volume x 100

The calculation of percentage of oxygen in the air after the evaporation of a volume of liquefied 

nitrogen or argon gas

The room dimensions



Volume of gas produced by evaporating liquid Argon

Note: 

This assumes that there is no dilution due to fresh air being pumped into the room and does not allow for the volume 

of any furniture or people in the room

Instructions:

 Fill in the white boxes with the appropriate figures and read the percentage of oxygen remaining in the room in 

the red box.

Percentage oxygen in room for ARGON evaporation

Percentage oxygen in room for NITROGEN evaporation

Room volume

Volume of gas produced by evaporating liquid Nitrogen

liquefied gas

		The calculation of percentage of oxygen in the air after the evaporation of a volume of liquefied nitrogen or argon gas

		Room volume

		Length (metres)		0.0		The room dimensions

		Width (metres)		0.0

		Height (metres)		0.0

		Room Volume (l*w*h)		0.0



		Volume of gas produced by evaporating liquid Nitrogen

		Volume of Liquid Nitrogen (litres)		0.00		The volume of liquid NITROGEN involved

		Volume of gas produced (Voll*682.5) (litres)		0.00		[liquid volume] times [expansion factor] (ref  data sheet)

		Volume of gas produced (metre3)(Volgas/1000)		0.000		volume converted to cubic metres



		Volume of gas produced by evaporating liquid Argon

		Volume of Liquid Argon (litres)		0.00		The volume of liquid ARGON involved

		Volume of gas produced (Vol*824.3) (litres)		0.00		[liquid volume] times [expansion factor] (ref data sheet)

		Volume of gas produced (metre3)(Volgas/1000)		0.000		volume converted to cubic metres



		Percentage oxygen in room for NITROGEN evaporation

		Vol0 0.2095([Room Vol]-[Vol gas produced]		0.00		Volume of oxygen = [fraction of oxygen in air] times ([the volume of the room] - [ volume of nitrogen liberated])

		%O2 (100*[VolO]/[Room vol]		ERROR:#DIV/0!		Volume of oxygen divided by the room volume x 100



		Percentage oxygen in room for ARGON evaporation

		Vol0 0.2095([Room Vol]-[Vol gas produced]		0.00		Volume of oxygen = [fraction of oxygen in air] times ([the volume of the room] - [ volume of argon liberated])

		%O2 (100*[VolO]/[Room vol]		ERROR:#DIV/0!		Volume of oxygen divided by the room volume x 100

		Instructions: Fill in the white boxes with the appropriate figures and read the percentage of oxygen remaining in the room in the red box.



		Note: This assumes that there is no dilution due to fresh air being pumped into the room and does not allow for the volume of any furniture or people in the room







cylinder gas

		The calculation of percentage of oxygen in the air after release of cylinder gas

		Room volume

		Length (metres)		0.0		The room dimensions

		Width (metres)		0.0

		Height (metres)		0.0

		Room Volume (l*w*h)		0.0



		Volume of gas released 

		Volume of gas [Litres]		0.00		cylinder contents at NTP

		Volume of gas produced (Metre3)   (Volgas/1000)		0.000		volume converted to cubic metres



		Percentage oxygen in room

		Vol0 0.2095([Room Vol]-[Vol gas produced]		0.00		Volume of oxygen = [fraction of oxygen in air] times ([the volume of the room] - [ volume of nitrogen liberated])

		%O2 (100*[VolO]/[Room vol]		ERROR:#DIV/0!		Volume of oxygen divided by the room volume times 100

		Fill in the white boxes and read the percentage of oxygen from the red box. Assumptions  -   even mixing of gas in the room and no allowances for fittings and persons in the room.

						ref:   based on a calculator by BOC






